## Slide 1: WOPM2024: OPM tutorial From sensors to a sensor array
Lauri Parkkonen

## Slide 2: SQUID sensors are far from the brain…
MEG with SQUID sensors
=> non-adaptable sensor helmet
=> large distance to brain
=> limited spatial resolution
Distance to brain (cm)
Iivanainen, Stenroos, Parkkonen (HBM 2013)

## Slide 3: If the sensors could be on the scalp…
Distance to brain (cm)
Distance to brain surface would reduce to less than half in adults. Even larger change in children.
Iivanainen, Stenroos, Parkkonen (HBM 2013)
On-scalp sensors
SQUIDs

## Slide 4
Noise level
SQUID
OPM (SERF)
Typ. 2–3 fT/rHz
Typ. 8–14 fT/rHz
Bandwidth
DC – several kHz
DC – ~150 Hz
Max. field
~10–20 nT change
1–2 nT absolute
Distance to scalp
20–30 mm; rigid array
4–9 mm; adjustable array
Temperature
Typ. –269 C
Typ. +150 C

## Slide 5: Comparing SQUID and OPM magnetometer arrays by simulations
State-of-the-art SQUID array optimally positioned
OPMs on the scalp
Virtual measurements based on MRIs of 10 adults
Arrays of 102 magnetometers

## Slide 6: OPM array increases amplitude
4
1
102-channel OPM vs. SQUID array:
Relative amplitude of neural sources in the cortex
Iivanainen, Stenroos, Parkkonen (Neuroimage 2017)
Ratio
OPMs better
SQUIDs better
3
2
Relative amplitude (ratio of the norms of the signal vectors). Average over 10 adult subjects.

## Slide 7: Somatosensory evoked responses
Electric stimulation of the left median nerve at wrist
Zetter, Iivanainen, Parkkonen (Sci Rep 2019)

## Slide 8: OPM array increases spatial resolution
Iivanainen, Stenroos, Parkkonen (Neuroimage 2017)
cm
Point-spread function of L2 minimum-norm estimate (thresholded at 50% of max value) of dipolar cortical sources

## Slide 9: Quantifying captured spatial information
OPM array
SQUID array
Spatial-frequency basis
Wavenumber / Wavelength
Component index
Field of a single dipole
Iivanainen, Mäkinen, Zetter, Stenroos, Ilmoniemi, Parkkonen; Neuroimage 2021

## Slide 10: OPM array increases spatial resolution
OPM vs. SQUID array; relative spatial resolution with ~100 channels
On-scalp (OPM) better
Ratio
Iivanainen, Stenroos, Parkkonen (Neuroimage 2017)

## Slide 11: OPM MEG: Towards non-invasive corticography
Clinical electrocorticographic (ECoG) grid: 3-mm diameter electrodes with 10-mm spacing
OPM MEG: Increase in spatial resolution should allow separating neural sources with a resolution comparable to ECoG measurements
Proper non-invasive detection of gamma-band (> 30 Hz) activity with OPM MEG?
Roland et al., 2013

## Slide 12: Vascular vs. electrophysiological signals
Mutual information between fMRI blood oxygenation signal and intracortical EEG amplitude envelopes
Magri et al. (J Neurosci, 2012)
Time (s)
Mutual info (bits)
Alpha 8–12 Hz
Beta 18–30 Hz
Gamma 40–100 Hz
LFP total 0–100 Hz
MUA 900–3000 Hz
Four monkeys
Visual stimulation and rest
4.7 T Bruker MRI
Multichannel electrodes in V1 and V2

## Slide 13
SQUID
OPM (SERF)
Noise level
Typ. 2–3 fT/rHz
Typ. 8–14 fT/rHz
Bandwidth
DC – several kHz
DC – ~150 Hz
Max. field
~10–20 nT change
1–2 nT absolute
Distance to scalp
20–30 mm; rigid array
4–9 mm; adjustable array
Temperature
Typ. –269 C
Typ. +150 C

## Slide 14: SQUID vs. OPM: Signal-to-noise ratio
SQUID sensor noise: 3 fT/rHz
OPM sensor noise: 9 fT/rHz
Iivanainen & Parkkonen, unpubl.
2
1
SNR ratio
OPM SNR
larger
SQUID SNR
larger
0
Count of cortical source points

## Slide 15: SQUID vs. OPM: Information content
SNR and spatial resolution are both captured in information-based metrics
OPM noise level is crucial!
Same amount of information does not mean the same information!
Iivanainen, Mäkinen, Zetter, Stenroos, Ilmoniemi, Parkkonen; Neuroimage 2021
of 3 fT/rHz

## Slide 16: Gamma-band responses to visual stimuli
Dynamic visual stimulus
9 OPM sensors (QZFM by QuSpin Inc.) above the occipital cortex
100 trials, total time ~10 min
Reference data from 306-ch. SQUID MEG (MEGIN Oy)
Iivanainen, Zetter, Parkkonen (Human Brain Mapping 2020)

## Slide 17: Gamma-band responses to visual stimuli
OPM magnetometer
SQUID magnetometer
Stimulus on
Iivanainen, Zetter, Parkkonen (Human Brain Mapping 2020)
OPM magnetometer

## Slide 18
SQUID
OPM (SERF)
Noise level
Typ. 2–3 fT/rHz
Typ. 8–14 fT/rHz
Bandwidth
DC – several kHz
DC – ~150 Hz
Max. field
~10–20 nT change
1–2 nT absolute (in open-loop mode)
Distance to scalp
20–30 mm; rigid array
4–9 mm; adjustable array
Temperature
Typ. –269 C
Typ. +150 C

## Slide 19: Compensation of residual field inside a magnetically shielded room
Iivanainen et al. (NeuroImage 2019)

## Slide 20: More subject-friendly compensation coils
https://bfieldtools.github.io/
Holmes et al. (Neuroimage 2018)
Mäkinen et al. (J Appl Phys 2020)
Zetter et al. (J Appl Phys 2020)
Coil windings restricted to two parallel planes
Coils optimized to accurately control fields and gradients in the measurement volume
Magnetic material distorts the coil fields => keep the coils away of shielded room walls
Low-noise current driver required

## Slide 21: Summary
Due to the proximity to brain, OPM arrays
  yield 2–4 times higher signals than SQUIDs for cortical sources
  capture higher spatial frequencies enabling higher spatial resolution
  produce “tighter” field patterns => proper placement of sensors important
OPM noise levels are ~3 times worse than those of SQUIDs
  Signal-to-noise ratio higher for superficial brain sources but worse for deeper brain sources
SERF OPMs require very low absolute magnetic field while SQUIDs do not
  Residual field compensation often needed/beneficial for OPMs

## Slide 22
+ many more people not in the picture